## Criteria
Category | Score | Description of Severity
Catastrophic | 5 | - Enforcement action by Regulatory agency such as consent decree, product seizure
- Very significant and catastrophic impact, significant losses and inefficiencies, necessitating immediate attention. Loss of operating capability, deterioration of efficiency
- Critical deviation from GMP requirements
- Batch failure : Batch does not or will not meet the specifications e.g. Any degradation product exceeding the specification, Contamination (microbial or other type), etc.
- Global product recalls (if any)
Critical | 4 | - Regulatory agency inspection may conclude serious non-compliance
- Likely product recalls from one or more markets
Moderate | 3 | - Regulatory agency inspection may identify a major concern, but deficiency quite easily resolved
- Limited product recalls possible
- Impact on operations and efficiency, but not pervasive. Management intervention required.
- Noticeable impact to product quality
Minor | 2 | - May require Regulatory agency notification
- Decision to release product not compromised
Negligible | 1 | - No regulatory issue
- No effect on and not noticeable by patient
- Limited or no impact on operations and quality of operational efficiency
- No impact to product quality and process robustness
Category | Score | Probability of Occurrence | Description Probability of Occurrence
Rare | 1 | <10% | Seen every 10 – 30 years
Unlikely | 2 | 10 – 24% | Seen every 5 – 9 years
Possible | 3 | 25 – 59% | Seen every 1 – 4 years
Likely | 4 | 60 -89% | Seen to occur 2 - 3 times a year
Almost certain | 5 | > 90% | Seen to occur more than 3 times a year
Category | Score | Description of Detectability
Undetectable | 5 | - No control exists
- Detection controls are absent
- Low likelihood that controls will detect the failure mode or its effects.
- Insufficient information to adequately assess and rate the control
Poor | 4 | - Very little chance of detection
- No procedural system in place, no risk reduction process procedure in place.
Moderate | 3 | - Not effective detection
- Likely to be detected
- Medium likelihood that controls will detect the failure mode or its effects.
- Mitigation plans or risk reduction process and procedures though in place but do not ensure adequate control over risk occurrence / impact as the risk still exists over its acceptance level.
Good | 2 | - Very likely to be detected
- Approved mitigation plans or risk reduction procedures, Checklists are laid down for early detection and immediate corrective actions are taken.
Excellent | 1 | - Very effective detection
- 100% likelihood of detection
- High likelihood that controls will detect the failure mode or its effects.
- Mitigation plans or risk reduction process and procedures involve stringent approval and reporting norms with responsibility duly mapped to various management levels, minimizing possibility of occurrence / optimizing protection in the case of occurrence of the same at very early stage, controlling and reducing the risk.

## FMEA
No. | Module ID | Risk Scenario (Detailed Description) | Impact | S | Cause | Control | P | D | RPN | Class | Mitigation Control | Test Ref (MT/InT/FT/RT) | S (Residual) | P (Residual) | D (Residual) | RPN (Residual) | Class (Residual)
1 | FRA-01 | Calculation Error: The system might miscalculate the duration of a temperature excursion. For instance, if a room stays above the limit for 20 minutes but the system counts it as 25 or more, it flags a false violation. | Data integrity loss: Incorrect excursion reporting leads to unnecessary investigations. | 3 | Logic error in the mathematical subtraction between timestamps. | Manual check or manual audit of rows. | 1 | 4 | 12 | MED* | Enforce strict 6-point logic for 25-minute threshold rules. | FT: OQ-TC-04, 06, 07; RT: PQ-TC-02 | 3 | 1 | 1 | 3 | LOW
2 | FRA-02 | Continuity Error: When data points are missing due to sensor downtime, the system might "connect the dots" across the gap, creating a fake violation period that didn't happen in reality. | Inaccurate violation periods shown on the report. | 3 | Large time gaps exist in the raw source CSV sensor data. | Manual audit of source files vs report. | 1 | 4 | 12 | MED* | Verify 5-minute intervals between every record; reset calculation if gap > 10 minutes. | FT: OQ-TC-05, 08; RT: PQ-TC-02 | 3 | 1 | 1 | 3 | LOW
3 | FRA-03 | Desync: When comparing pressure between two rooms, the system might accidentally compare Room A data from 10:00 AM with Corridor data from 10:05 AM, giving a wrong result. | False pressure hierarchy violations reported to the user. | 4 | High complexity in the data joining logic between multiple files. | None currently. | 2 | 4 | 32 | HIGH** | Use strict "Inner Join" on exact timestamps for all rooms to ensure the pressure difference is always calculated using data from the same moment. | FT: OQ-TC-10, 11, 12, 13 | 4 | 1 | 1 | 4 | LOW
4 | FRA-04 | Tamper Vulnerability: A user could manually open the audit log in Notepad and edit the text to hide a mistake or change the history of actions, breaking regulatory compliance (21 CFR Part 11). | Compliance failure: Electronic records are not secure or trustworthy. | 4 | Plain text log files are technically accessible on the Windows drive. | None. | 2 | 4 | 32 | HIGH** | Implement SHA-256 Hash-chaining for every log entry. If any letter is changed manually, the "chain" breaks and the system alerts the team. | MT: IQ-TC-07; FT: OQ-TC-23, 25; RT: PQ-TC-03 | 4 | 1 | 1 | 4 | LOW
6 | FRA-05 | Deployment Failure: The application might fail to open on a new computer because it cannot find its internal icons, graphics, or hidden files due to hardcoded paths. | System unavailability and delay in quality review processes. | 3 | Path resolution issues when software is moved to different Windows workstations. | Application crash is highly visible. | 2 | 2 | 12 | MED* | Use "Path-Independent" logic to automatically search internal system folders for assets regardless of workstation. | MT: IQ-TC-01, 04, 06; RT: PQ-TC-03 | 3 | 1 | 1 | 3 | LOW
7 | FRA-06 | Data Skewing: A sensor or network glitch might cause the same data row to be saved twice with the same timestamp, causing the system to double-count the time. | Inaccurate statistics; double-counting excursion duration. | 3 | Glitches in sensor or network transmission causing duplicate records. | Requires row-by-row manual check to identify. | 2 | 3 | 18 | MED* | Programmatically perform a "Deduplication" pass, automatically deleting any rows with duplicate timestamps before analysis. | RT: PQ-TC-02 | 3 | 1 | 1 | 3 | LOW
8 | FRA-07 | Data Destruction: A software bug could accidentally try to "Save" over the original raw sensor data files, corrupting or deleting the records required for audits. | Irrecoverable loss of original GxP data records required for audits. | 4 | Logic error in the file handling/saving part of the software code. | Visible only after the file is already lost. | 1 | 3 | 12 | HIGH** | Enforce strict "Read-Only" commands at the code level so the software cannot physically modify source files. | MT: IQ-TC-05 | 4 | 1 | 1 | 4 | LOW
10 | FRA-08 | Mapping Error: During file splitting, the system might get confused and accidentally move Room A's temperature data into Room B's file. | Incorrect data attribution: Data linked to wrong physical location. | 3 | Inconsistent room naming or formatting in the bulk source file. | Hard to detect if data values look similar. | 2 | 3 | 18 | MED* | Use intelligent "Regex" pattern recognition to securely lock data columns to their corresponding Room IDs. | FT: OQ-TC-14; RT: PQ-TC-04 | 3 | 1 | 1 | 3 | LOW
11 | FRA-09 | Data Loss: During processing of very large bulk files, the system might fail to capture all records, losing data from the end of the day. | Incomplete records leading to missing excursions or audit gaps. | 3 | Buffer overflows or premature "End of File" signals during processing. | Only visible if manually checking row counts. | 1 | 3 | 9 | MED* | Implement row count verification and specifically verify the presence of the mandatory "End of Report" footer text. | FT: OQ-TC-14; RT: PQ-TC-04 | 3 | 1 | 1 | 3 | LOW
12 | FRA-10 | Header Integrity: The newly split CSV file might have a corrupted or messy format that the analysis software cannot read, halting the review process. | Analysis engine fails to open the file, stopping the review process. | 3 | Logic error in the reconstruction of the CSV file structure. | Highly visible error during analysis step. | 1 | 3 | 9 | MED* | Standardize header reconstruction logic using a meticulously defined template to match original reports. | InT: OQ-TC-19; FT: OQ-TC-14; RT: PQ-TC-04 | 3 | 1 | 1 | 3 | LOW
*Minimum Medium due to S=3 rule. **High Risk due to S=4/5 rule.