Enclosure 2: Test Script

| Test ID | Test Scenario | Expected Result |
| --- | --- | --- |
| OQ-TC-02 | To confirm the system blocks analysis if the configuration file (`SetPointLimit.xlsx`) is missing. | - The system displays an error message: "ERR-002: Limit File Not Found". - The 'Analyze' functionality is disabled. |
| OQ-TC-03 | To verify the system handles non-numeric limit values in the Excel file without crashing. | - The system stops the process and displays: "ERR-003: Invalid Configuration - High Limit must be a number" |
| OQ-TC-19 | To verify that the system rejects files with incorrect or missing column headers. | - The system displays an error: "ERR-005: Invalid File Format - Required columns not found." |
| OQ-TC-24 | To confirm that every data file (CSV) analyzed by the system creates a unique digital fingerprint in the audit log. To trust the final report, we must prove it came from a specific, original file that has not been changed. A SHA-256 hash is like a "digital fingerprint." If the file changes by even one byte, the fingerprint changes completely. | - The log entry contains the full name of the processed file. - The log entry contains a unique 64-character SHA-256 hash string. |
| IQ-TC-06 | To mathematically prove that the application is fully independent. It must be capable of executing all its core functions without relying on external internet connections or pre-installed programming software (like Python) on the user's computer. | - The application launches successfully in a completely offline environment. - The visual dashboard is fully accessible and functional through the local network port. |
| OQ-TC-23 | To verify that the system detects unauthorized modifications to the `audit_trail.log` file and halts execution. According to regulatory rules (21 CFR Part 11), electronic records must be protected from manual changes. The system uses a "hash-chain" where each line of text is linked to the previous one mathematically. If even one letter is changed, the link breaks. This ensures that no one has edited the history of what the system did. | - The application must not open the main dashboard. - A popup window or console message appears with the exact text: "FATAL ERROR 004: Audit Trail Integrity Check Failed" |